## Arkusz1
CALIBRATION REPORT : K220/19
CLIENT | RADWAG Wagi Elektroniczne
ul. Toruńska 5 26-600 Radom
USER
CALIBRATION OBJECT - CHARACTERISTICS [B]
Object name | Reference weight |  |  |  | Material | stainless steel
S/N | 876665 |  |  |  | Material density | 8000 kg/m3
Manufacturer | RADWAG Wagi Elektroniczne |  |  |  | Material dens. Det. Uncert. | 140 kg/m3
Year of manufacture | 2018 |  |  |  | Shape | cylindrical form
Marking | 5R |  |  |  | Maximum Permissible Error | 0.000016 g
Class (OIML R-111) | E1 |  |  |  | Air Density | 1.2 kg/m3
Nominal mass | 5 g |  |  |  | Set no | 876665
Adjustment | monobloc without adjusting hole
REFERENCE WEIGHT - CHARACTERISTICS [A]
S/N | G0601916 |  |  |  | Mass | 5.0000043 g
Manufacturer | RADWAG Wagi Elektroniczne |  |  |  | Calibration uncertainty | 0.000005 g
Manufacturer | 2006 |  |  |  | Calibration Certificate | 93479
Year of manufacture | R |  |  |  | Material density | 0 kg/m3
Class (OIML R-111) | E1 |  |  |  | Material dens. Det. Uncert. | 0 kg/m3
REFERENCE WEIGHT - CHARACTERISTICS [A]
Marking/type | UMA 5 |  |  |  | Scale interval | 0.0000001 g
S/N | 61800 |  |  |  | Standard deviation | 0.000005 g
Manufacturer | RADWAG Wagi Elektroniczne |  |  |  | Year of manufacture | 2019
MEASURE DEVICE
Marking/type | Temperature [C] |  |  |  | Humidity [%] | Pressure [hPa]
S/N | 0.06 |  |  |  | -0.1 | -0.1
Manufacturer | 0.25 |  |  |  | 1.4 | 0.7
AMBIENT CONDITIONS AT MEASUREMENT START
| Temperature [C] |  |  |  | Humidity [%] | Pressure [hPa]
Measurement | 25.34 |  |  |  | 26.9 | 976.6
Measurement correction | 25.40 |  |  |  | 26.8 | 976.5
AMBIENT CONDITIONS AT MEASUREMENT END
| Temperature [C] |  |  |  | Humidity [%] | Pressure [hPa]
Measurement | 25.34 |  |  |  | 27 | 977
Measurement correction | 25.40 |  |  |  | 26.90 | 976.90
Air Density | 1.1361 kg/m3
Average DIFFERENCE (B-A): |  | 0,0000130 g
Standard deviation: |  | 0,00000037g
Measurement uncertainty: |  | 0,0000025g
Coverage factor: |  | 2
Calibration result: |  | 5,000173g
Calibration uncerainty |  | 5e-06
Calibration start time: |  | 2019-04-03
Calibration end time: |  | 2019-04-03 11:48
Calibration mass: |  | 0g
Calibration certificate
Order number: |  | Z7/19
CYCLE NO.1 | 1 |  |  | DIFFERENCE (B-A): 0,0000135 g
OBJECT |  |  | INDICATION
CONTROL
TEST
TEST
CONTROL
CYCLE NO.2 | 2 |  |  | DIFFERENCE (B-A): 0,0000135 g
OBJECT |  |  | INDICATION
CONTROL
TEST
TEST
CONTROL
CYCLE NO.3 | 3 |  |  | DIFFERENCE (B-A): 0,0000135 g
OBJECT |  |  | INDICATION
CONTROL
TEST
TEST
CONTROL
CYCLE NO.4 | 4 |  |  | DIFFERENCE (B-A): 0,0000135 g
OBJECT |  |  | INDICATION
CONTROL
TEST
TEST
CONTROL
CYCLE NO.5 | 5 |  |  | DIFFERENCE (B-A): 0,0000135 g
OBJECT |  |  | INDICATION
CONTROL
TEST
TEST
CONTROL
CYCLE NO.6 | 6 |  |  | DIFFERENCE (B-A): 0,0000135 g
OBJECT |  |  | INDICATION
CONTROL
TEST
TEST
CONTROL
CYCLE NO.7 | 7 |  |  | DIFFERENCE (B-A): 0,0000135 g
OBJECT |  |  | INDICATION
CONTROL
TEST
TEST
CONTROL